| Series | Magnification | Numerical App | WD(mm) | Coverslip (mm) | Notes | Calibration |
| --- | --- | --- | --- | --- | --- | --- |
| UPlan Fl | 4 | 0.13 | 17 | - | IX50 |  |
| UPlan Fl | 10 | 0.3 | 10 | - | IX50 |  |
| LcPlanFl | 20 | 0.4 | 6.9 | 0-2.5 | IX50 |  |
| LcPlanFl | 40 | 0.6 | 2.6 | 0-2.5 | IX50 |  |
| UPlanFL | 60 | 1.25 | 0.1 | 0.17 | AX70 (7.7p/µm) | 0.13µ/px |
| UPlanApo | 40 | 1.0 | .15 | 0.17 | AX70(5.12p/µm) | 0.20µ/px |
| UPlanFL | 20 | 0.5 | 1.6 | 0.17 | AX70(2.53p/µm) | 0.40µ/px |
| UPlanFL | 10 | 0.3 | 10 | - | AX70(1.28p/µm) | 0.78µ/px |
| LcPlanFl | 40 | 0.6 | 2.6 | 0-2.5 | IX70 |  |
| LcPlanFl | 60 | 0.7 | 1.7 | 0-2.5 | IX70 |  |
| UPlanFL | 40 | 0.75 | 0.51 | 0.17 | IX70 |  |
| UPlanFL | 10 | 0.3 | 10 | - | IX70 |  |
| UPlanFL | 10 | 0.3 | 10 | - | BX50 | 1.38 |
| UPlanFL | 20 | 0.5 | 1.6 | 0.17 | BX50 | 0.69 |
| UPlanApo | 40 | 1.0 | .13 | 0.11-.23 | BX50 | 0.345 |
| PlanApo | 60 | 1.4 | .15 | 0.17 | BX50 | 0.230 |
| LumplanFL | 40 | 0.8 | 3.3 | Water lens | BX50 | 0.345 |
| UmplanFL | 10 | 0.3 | 3.3 | Water lens | BX61 | 0.625 |
| LumplanFL | 40 | 0.8 | 3.3 | Water lens | BX61 | 0.16125 |
| LumplanFL | 60 | 0.9 | 2.0 | Water lens | BX61 | 0.1075 |
| UPlanFL | 40 | 1.3 | .5 |  | BX61 | 0.16125 |
| UPlanApo | 60 | 1.4 | 0.15 | 0.17 | BX61 | 0.1075 |
| PlanApo | 10 | 0.4 |  |  | BX61 | 0.625 |
| uPLANSAPO | 20 | .75 |  | 0.17 | BX61 | 0.3225 |
| Plan Fluor | 4 | 0.13 | 17.1 |  | NiE(.62p/µ) | 1.61µm/px |
| Plan Fluor | 10 | 0.3 | 16 |  | NiE(.1.56/µ) | 0.64µm/px |
| Plan Apo λ | 20 | 0.75 | 1 |  | NiE(3.13p/µ) | 0.32µm/px |
| Plan Apoλ | 40 | 0.95 | .11-.23 | Collar | NiE(6.25p/µ) | 0.16µm/px |
| Plan Apoλ | 10 | 0.45 | 4 |  | C2Si(@512) | 1.64µm/px |
| Plan Apo VC | 20 | .75 | 1 |  | C2Si | 0.82µm |
| Apo LWD | 40 | 1.15 | 0.2 | Water lens | C2Si | 0.41µm |
| Plan Apoλ | 60 | 1.4 | 0.14 | Oil | C2Si | 0.273µm |

For the Fluoview 1:

|  | Resolution captured |  |  |
| --- | --- | --- | --- |
| Mag | 1024x1024 | 1024x768 | 800x600 |
| 10x | 1.38 µm/pixel | 1.56 µm/pixel | 2 µm/pixel |
| 20x | 0.69 µm/pixel | 0.78 µm/pixel | 1 µm/pixel |
| 40x | 0.345 µm/pixel | 0.39 µm/pixel | 0.5 µm/pixel |
| 60x | 0.23 µm/pixel | 0.26 µm/pixel | 0.33 µm/pixel |
| 100x | 0.138 µm/pixel | 0.156 µm/pixel | 0.2 µm/pixel |

For the Olympus DSU:

| Resolution captured |  |  |
| --- | --- | --- |
|  | 1344x1024 |  |
| 10 | 0.625 µm/px |  |
| 20 | 0.32µm/px |  |
| 40 | 0.16µm/px |  |
| 60 | 0.1075µm/px |  |

For the Fluoview 10i:

| Resolution captured |  |  |
| --- | --- | --- |
| Mag | 1024x1024 | 512x512 |
| Overview | 3584 pixels=8909.6µm (7x7 512x512) |  |
| 10x | 1.244 | 2.488 |
| 60x | 0.207 | 0.415 |

DEALING WITH ZOOM: Base calibration ÷zoom factor=new calibration

For example, using the 10x with a zoom of 2 (20x total mag) @1024

1.244µ/pixel divided by zoom factor of 2 = 0.622 µm per pixel at 20x @1024

For the Nikon NiE:

| Mag | Qi1Mc |  |  | Fi2 |  |
| --- | --- | --- | --- | --- | --- |
|  | 1280x1024 | 640x512 | 2560x1920 |  | 1280x960 |
| 4x | 1.61µm/px | 3.23 µm/pixel | 0.85 µm/pixel |  | 1.7 µm/pixel |
| 10x | 0.64µm/px | 1.29 µm/pixel | 0.34 µm/pixel |  | 0.68 µm/pixel |
| 20x | 0.32µm/px | 0.64 µm/pixel | 0.17 µm/pixel |  | 0.34 µm/pixel |
| 40x | 0.16µm/px | 0.32 µm/pixel | 0.085 µm/pixel |  | 0.17 µm/pixel |

For the C2Si: 7/31/2018 cal data

| Resolution captured |  |  |
| --- | --- | --- |
| Mag | 1024x1024 | 512x512 |
| 10x | 0.819 | 1.638 |
| 20x | 0.41 | 0.82 |
| 40x (w) | 0.21 | 0.41 |
| 60x (o) | 0.137 | 0.273 |

DEALING WITH ZOOM: Base calibration ÷zoom factor=new calibration

For example, using the 10x with a zoom of 2 (20x total mag) @1024

0.819/pixel divided by zoom factor of 2 = 0.41 µm per pixel at 20x @1024

*****For the Nikon C2Si:unknown source data

| Mag | Resolution captured |  |  |
| --- | --- | --- | --- |
|  | 2048x2048 | 1024x1024 | 512x512 |
| 10x | 0.62µm/pixel | 1.24µm/pixel | 2.47µm/pixel |
| 20x | 0.31µm/pixel | 0.62µm/pixel | 1.24µm/pixel |
| 40x | 0.16µm/pixel | 0.31µm/pixel | 0.62µm/pixel |
| 60x | 0.1µm/pixel | 0.21µm/pixel | 0.41µm/pixel |

| NOT in use |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Plan NeoFl | 10 | 0.3 | 5.6 |  | ZEISS |  |
| AcroPlan | 40 | 0.8 | 3.6 | Water lens | ZEISS |  |
| Plan Apo | 63 | 1.4 | 0.18 | 0.17 | ZEISS |  |